Tabular Summary of selected relevant points from ISO/IEC 17025:2005

| No. | Title of the clause | Extracts from text of 17025:2005 and key words (arranged in blocks) | Comment |
| --- | --- | --- | --- |
|  | Introduction | Accreditation bodies that recognize the competence of testing and calibration laboratories should use this International Standard as the basis for their accreditation. Laboratories should have a quality management system that conforms to the principles of ISO 9001:2000 (which has superseded ISO 9001:1994 and ISO 9002:1994). ISO 17025:2005 incorporates all of those requirements of ISO 9001:2000 that are relevant to the scope of testing and calibration services. If a laboratory wishes accreditation to ISO 17025:2005 it should select an accreditation body that operates in accordance with ISO 17011. | Selected staffs should possess and work through a personal copy of ISO/IEC 17025:2005 ISO/IEC 17000 ISO/IEC 17011 ISO9001:2000 VIM, International vocabulary of basic and general terms in metrology |
| 1 | Scope | ISO 17025:2005 specifies the general requirements a laboratory has to meet if it is to be recognized as competent to carry out tests and/or calibrations, including sampling. It covers testing and calibration performed using standardized methods, methods not covered by standardized methods, and laboratory-developed methods. It is applicable to all organizations performing tests and/or calibrations regardless of the number of personnel or the extent of the scope of testing and/or calibration activities. Compliance with regulatory and safety requirements on the operation of laboratories is not covered by ISO 17025:2005. |  |
| 2 | Normative references | ISO 17000, Conformity assessment — Vocabulary and general principles. VIM, International vocabulary of basic and general terms in metrology, issued by BIPM, IEC, IFCC, ISO, IUPAC, IUPAP and OIML |  |
| 3 | Terms and definitions | The relevant terms and definitions given in ISO 17000 and VIM apply. |  |
| 4 | Management requirements | Clause 4 specifies the requirements for sound management. |  |
| 4.1 | Organization | The laboratory or the organization of which it is part shall be an entity that can be held legally responsible. The organization and management structure of the laboratory, its place in any parent organization, and the relationships between quality management, technical operations and support services shall be defined; The managerial and technical personnel of the laboratory which has the overall responsibility for the technical operations shall have the authority and resources needed to carry out their duties, including the implementation, maintenance and improvement of the management system, to identify departures from the management system or from the procedures for performing tests and/or calibrations, and to initiate actions to prevent or minimize such departures. The responsibility, authority and interrelationships of all personnel who manage, perform or verify work affecting the quality of the tests and/or calibrations shall be specified; Top management shall appoint a member of staff as quality manager who shall have defined responsibility and authority for ensuring that the management system related to quality is implemented and followed at all times; the quality manager shall have direct access to the highest level of management at which decisions are made on laboratory policy or resources; deputies for key managerial personnel shall be appointed; Management shall provide adequate supervision of testing and calibration staff, by persons familiar with methods and procedures, purpose of each test and/or calibration, and with the assessment of the test or calibration results; Management shall ensure appropriate communication processes within the laboratory. Personnel shall be aware of the relevance and importance of their activities and how they contribute to the achievement of the objectives of the management system and communication regarding the effectiveness of the management system shall take place. The management and personnel of the laboratory shall be free from any undue internal and external pressures and influences that may adversely affect the quality of their work; There shall be policies and procedures to avoid involvement in any activities that would diminish confidence in its competence, impartiality, judgement or operational integrity; it shall have policies and procedures to ensure the protection of its customers' confidential information and proprietary rights, including procedures for protecting the electronic storage and transmission of results; |  |

| No. | Title of the clause | Extracts from text of 17025:2005 and key words (arranged in blocks) | Comment |
| --- | --- | --- | --- |
| 4.2 | Management system | The laboratory shall establish, implement and maintain a management system appropriate to the scope of its activities; the laboratory shall document its policies, systems, programmes, procedures and instructions to the extent necessary to assure the quality of the test and/or calibration results. The system’s documentation shall be communicated to, understood by, available to, and implemented by the appropriate personnel. The laboratory's management system policies related to quality shall be defined in a quality manual. It should contain: A statement of the overall objectives which shall be reviewed during management review; A quality policy statement issued under the authority of top management which includes in a concise document at least: the management's commitment to good professional practice, to the quality of its testing and calibration, and to compliance with the International Standard; a management statement of the laboratory's standard of service; the purpose of the management system related to quality; a management commitment to continually improve the effectiveness of the management system; evidence of commitment to the development and implementation of the management system and to continually improving its effectiveness; assurance that the integrity of the management system is maintained when changes to the management system are planned and implemented; the roles and responsibilities of technical management and the quality manager, including their responsibility for ensuring compliance with ISO 17025:2005; a requirement that all personnel concerned with testing and calibration activities within the laboratory familiarize themselves with the quality documentation and implement the policies and procedures in their work; communication to the organization the importance of meeting customer requirements as well as statutory and regulatory requirements; documentation or reference to the supporting procedures including technical procedures and an outline of the structure of the documentation used in the management system. |  |
| 4.3 | Document control |  |  |
| 4.3.1 | General | The laboratory shall establish and maintain procedures to control all documents that form part of its management system (internally generated or from external sources), such as regulations, standards, other normative documents, test and/or calibration methods, as well as drawings, software, specifications, instructions and manuals. NOTE: In this context “document” could be policy statements, procedures, specifications, calibration tables, charts, text books, posters, notices, memoranda, software, drawings, plans, etc. These may be on various media, whether hard copy or electronic, and they may be digital, analog, photographic or written. The control of data related to testing and calibration is covered in 5.4.7. The control of records is covered in 4.13. |  |
| 4.3.2 | Document approval and issue | All documents issued to personnel in the laboratory as part of the management system shall be reviewed and approved for use by authorized personnel prior to issue and shall be periodically reviewed and, where necessary, revised; authorized editions of appropriate documents shall be available at all locations where operations essential to the effective functioning of the laboratory are performed; invalid or obsolete documents shall be promptly removed from all points of issue or use; Management system documents generated by the laboratory shall be uniquely identified. Such identification shall include the date of issue and/or revision identification, page numbering, the total number of pages or a mark to signify the end of the document, and the issuing authority(ies). A master list or an equivalent document control procedure identifying the current revision status and distribution of documents in the management system shall be established and shall be readily available to preclude the use of invalid and/or obsolete documents. |  |
| 4.3.3 | Document changes | Changes to documents shall be reviewed and approved by the same function that performed the original review unless specifically designated otherwise. Procedures shall be established to describe how changes in documents are made and controlled. A revised document shall be formally re-issued as soon as practicable. |  |

| No. | Title of the clause | Extracts from text of 17025:2005 and key words (arranged in blocks) | Comment |
| --- | --- | --- | --- |
| 4.4 | Review of requests, tenders and contracts | The laboratory shall have policies and establish and maintain procedures for the review of requests, tenders and ensuring that the requirements, including the methods to be used, are adequately defined, documented and understood, the laboratory has the capability and resources to meet the requirements, the appropriate test and/or calibration method is selected and is capable of meeting the customers' requirements. The review shall also cover any work that is subcontracted by the laboratory. If a contract needs to be amended after work has commenced, the same contract review process shall be repeated and any amendments shall be communicated to all affected personnel. The customer shall be informed of any deviation from the contract. Records of reviews, including any significant changes, shall be maintained. | Before going into more details here, it should be decided whether this clause is relevant for the QC laboratories and who could be “customers” in such cases. |
| 4.5 | Subcontracting of tests and calibrations | When a laboratory subcontracts work this work shall be placed with a competent subcontractor (e.g., who complies with ISO 17025:2005). The laboratory shall advise the customer of the arrangement in writing and is responsible to the customer for the subcontractor’s work. The laboratory shall maintain a register of all subcontractors that it uses for tests and/or calibrations and a record of the evidence of compliance with ISO 17025:2005 for the work in question. |  |
| 4.6 | Purchasing services and supplies | The laboratory shall have a policy and procedure(s) for the selection and purchasing of services and supplies and the storage of reagents and laboratory consumable materials relevant for the quality of tests and calibrations. Such supplies, reagents and consumable materials shall not be used until they have been verified as complying with standard specifications and requirements defined in the methods for the tests and/or calibrations concerned. Records of compliance checks shall be maintained. Suppliers of critical consumables, supplies and services shall be evaluated and records of these evaluations and a list of those approved shall be maintained. |  |
| 4.7 | Service to the customer | The laboratory shall cooperate with customers in clarifying the customer's request and in monitoring the laboratory’s performance in relation to the work performed, e.g., by providing the customer reasonable access to relevant areas of the laboratory for the witnessing of tests and/or calibrations performed for the customer and by providing test and/or calibration items needed by the customer for verification purposes. The laboratory should inform the customer of any delays or major deviations in the performance of the tests and/or calibrations. The laboratory shall seek feedback from its customers (e.g., satisfaction surveys) and analyse the feedback to improve the management system, testing and calibration activities and customer service. |  |
| 4.8 | Complaints | The laboratory shall have a policy and procedure for the resolution of complaints received from customers or other parties. Records shall be maintained of all complaints and of the investigations and corrective actions taken by the laboratory. |  |
| 4.9 | Control of nonconforming testing and/or calibration work | The laboratory shall have a policy and procedures that shall be implemented when any aspect of its testing and/or calibration work, or the results of this work, do not conform to its own procedures or the agreed requirements of the customer. They shall ensure that responsibilities and authorities for the management of nonconforming work (evaluation of the significance, corrective actions (see 4.11), notification of customer and recall of work – where necessary – and resumption of work) are designated. NOTE: Identification of nonconforming work or problems with the management system or with testing and/or calibration activities can occur at various places within the management system and technical operations. Examples are customer complaints, quality control, instrument calibration, checking of consumable materials, staff observations or supervision, test report and calibration certificate checking, management reviews and internal or external audits. |  |
| 4.10 | Improvement | The laboratory shall continually improve the effectiveness of its management system through the use of the quality policy, quality objectives, audit results, analysis of data, corrective and preventive actions and management review. |  |
| 4.11 | Corrective action |  |  |
| 4.11.1 | General | The laboratory shall establish a policy and a procedure and shall designate appropriate authorities for implementing corrective action when nonconforming work or departures from the policies and procedures in the management system or technical operations have been identified. |  |
| 4.11.2 | Cause analysis | The procedure for corrective action shall start with an investigation to determine the root cause(s) of the problem. |  |
| 4.11.3 | Selection and implementation of corrective actions | The laboratory shall identify potential corrective actions appropriate to the magnitude and the risk of the problem and select and implement the action(s) most likely to eliminate the problem and to prevent recurrence. |  |
| No. | Title of the clause | Extracts from text of 17025:2005 and key words (arranged in blocks) | Comment |
| 4.11.4 | Monitoring of corrective actions | The laboratory shall monitor the results to ensure that the corrective actions taken have been effective. |  |
| 4.11.5 | Additional audits | Where the identification of nonconformities or departures casts (serious) doubts on the laboratory's compliance with its own policies and procedures, or on its compliance with ISO 17025:2005, the laboratory shall audit the appropriate areas of activity in accordance with 4.14 as soon as possible. |  |
| 4.12 | Preventive action | Needed improvements and potential sources of technical or management system nonconformities shall be identified and action plans shall be developed, implemented and their effectiveness monitored to take advantage of the opportunities for improvement and to reduce the likelihood of the occurrence of nonconformities. |  |
| 4.13 | Control of records |  |  |
| 4.13.1 | General | The laboratory shall establish and maintain procedures for identification, collection, indexing, access, filing, storage, maintenance and disposal of quality and technical records. Quality records shall include reports from internal audits and management reviews as well as records of corrective and preventive actions. All records shall be held secure and in confidence; the laboratory shall have procedures to protect and back-up records stored electronically and to prevent unauthorized access to or amendment of these records; Records shall be legible and shall be stored and retained in such a way that they are readily retrievable in facilities that provide a suitable environment to prevent damage or deterioration and to prevent loss. Retention times of records shall be established. |  |
| 4.13.2 | Technical records | The laboratory shall retain – as far as possible - records of original observations, derived data and sufficient information to establish an audit trail, calibration records, staff records and a copy of each test report or calibration certificate issued, for a defined period. The records for each test or calibration shall contain sufficient information to facilitate, if possible, identification of factors affecting the uncertainty and to enable the test or calibration to be repeated under conditions as close as possible to the original. Observations, data and calculations shall be recorded at the time they are made and shall be identifiable to the specific task. The records shall include the identity of personnel responsible for the sampling, performance of each test and/or calibration and checking of results. When mistakes occur in records, each mistake shall be crossed out, not erased, made illegible or deleted, and the correct value entered alongside. All such alterations to records shall be signed or initialled by the person making the correction. In the case of records stored electronically, equivalent measures shall be taken to avoid loss or change of original data. NOTE : Technical records are accumulations of data (see 5.4.7) and information which result from carrying out tests and/or calibrations and which indicate whether specified quality or process parameters are achieved. They may include forms, contracts, work sheets, work books, check sheets, work notes, control graphs, external and internal test reports and calibration certificates, customers' notes, papers and feedback. |  |
| 4.14 | Internal audits | The laboratory shall periodically (normally a cycle of one year), and in accordance with a predetermined schedule and procedure, conduct internal audits of its activities to verify that its operations continue to comply with the requirements of the management system and ISO 17025. The internal audit programme shall address all elements of the management system, including the testing and/or calibration activities. It is the responsibility of the quality manager to plan and organize audits as required by the schedule and requested by management. Such audits shall be carried out by trained and qualified personnel who are, wherever resources permit, independent of the activity to be audited. When audit findings cast doubt on the effectiveness of the operations or on the correctness or validity of the laboratory's test or calibration results, the laboratory shall take timely corrective action, and shall notify customers in writing if investigations show that the laboratory results may have been affected. The area of activity audited, the audit findings and corrective actions that arise from them shall be recorded. Follow-up audit activities shall verify and record the implementation and effectiveness of the corrective action taken. |  |

| No. | Title of the clause | Extracts from text of 17025:2005 and key words (arranged in blocks) | Comment |
| --- | --- | --- | --- |
| 4.15 | Management reviews | In accordance with a predetermined schedule (typically once every 12 months) and procedure, the laboratory’s top management shall periodically conduct a review of the laboratory's management system and testing and/or calibration activities to ensure their continuing suitability and effectiveness, and to introduce necessary changes or improvements. Findings from management reviews and the actions that arise from them shall be recorded. The management shall ensure that those actions are carried out within an appropriate and agreed timescale. The review shall take account of: the suitability of policies and procedures; reports from managerial and supervisory personnel; the outcome of recent internal audits; corrective and preventive actions; assessments by external bodies; the results of interlaboratory comparisons or proficiency tests; changes in the volume and type of the work; customer feedback; complaints; recommendations for improvement; other relevant factors, such as quality control activities, resources and staff training. NOTE: Results should be considered at regular management meetings and should feed into the laboratory planning system. |  |
| 5 | Technical requirements | Clause 5 specifies the requirements for technical competence for the type of tests and/or calibrations the laboratory undertakes. |  |
| 5.1 | General | Many factors determine the correctness and reliability of the tests and/or calibrations performed by a laboratory (see 5.2 to 5.8). The extent to which the factors contribute to the total uncertainty of measurement differs considerably between (types of) tests and between (types of) calibrations. The laboratory shall take account of these factors in developing test and calibration methods and procedures, in the training and qualification of personnel, and in the selection and calibration of the equipment it uses. |  |
| 5.2 | Personnel | The laboratory management shall ensure the competence of all who operate specific equipment, perform tests and/or calibrations, evaluate results, and sign test reports and calibration certificates. Personnel performing specific tasks shall be qualified on the basis of appropriate education, training, experience and/or demonstrated skills, as required. Personnel who are employed by, or under contract to, the laboratory shall be used. The laboratory shall maintain current job descriptions for managerial, technical and key support personnel involved in tests and/or calibrations. The laboratory shall ensure that personnel are supervised. The management shall authorize specific personnel to perform particular types of sampling, test and/or calibration, to issue test reports and calibration certificates, to give opinions and interpretations and to operate particular types of equipment. The laboratory shall maintain records of the relevant authorization(s), competence, educational and professional qualifications, training, skills and experience of all technical personnel, including contracted personnel. This information shall be readily available and shall include the date on which authorization and/or competence is confirmed. The management of the laboratory shall formulate the goals with respect to the education, training and skills of the laboratory personnel. The laboratory shall have a policy and procedures for identifying training needs and providing training of personnel. The training programme shall be relevant to the present and anticipated tasks of the laboratory. The effectiveness of the training actions taken shall be evaluated. NOTE: In some technical areas (e.g. non-destructive testing) it may be required that the personnel performing certain tasks hold personnel certification. NOTE: Job descriptions should be define (as a minimum) the responsibilities with respect to planning, performing, and evaluation of results of tests and/or calibrations; responsibilities with respect to method modification and development and validation of new methods; responsibilities for reporting opinions and interpretations. |  |
| 5.3 | Accommodation and environmental conditions | Laboratory facilities for testing and/or calibration shall be such as to facilitate correct performance of the tests and/or calibrations and environmental conditions shall not adversely affect the required quality of any measurement. The laboratory shall monitor, control and record environmental conditions as required by the relevant specifications, methods and procedures or where they influence the quality of the results. Due attention shall be paid, for example, to biological sterility, dust, electromagnetic disturbances, radiation, humidity, electrical supply, temperature, and sound and vibration levels, as appropriate. Tests and calibrations shall be stopped when the environmental conditions jeopardize the results. There shall be effective separation between neighbouring areas in which there are incompatible activities. Access to and use of areas affecting the quality of the tests and/or calibrations shall be controlled. |  |
| No. | Title of the clause | Extracts from text of 17025:2005 and key words (arranged in blocks) | Comment |
| 5.4 | Test and calibration methods and method validation |  |  |
| 5.4.1 | General | The laboratory shall use appropriate methods and procedures for all tests and/or calibrations within its scope. These include sampling, handling, transport, storage and preparation of items to be tested and/or calibrated, and, where appropriate, an estimation of the measurement uncertainty as well as statistical techniques for analysis of test and/or calibration data. The laboratory shall have instructions on the use and operation of all relevant equipment, and on the handling and preparation of items for testing and/or calibration, or both, where the absence of such instructions could jeopardize the results of tests and/or calibrations. All instructions, standards, manuals and reference data relevant to the work of the laboratory shall be kept up to date and shall be made readily available to personnel (see 4.3). Deviation from test and calibration methods shall occur only if the deviation has been documented, technically justified, authorized, and accepted by the customer. NOTE International, regional or national standards or other recognized specifications that contain sufficient and concise information on how to perform the tests and/or calibrations do not need to be supplemented or rewritten as internal procedures if these standards are written in a way that they can be used as published by the operating staff in a laboratory. It may be necessary to provide additional documentation for optional steps in the method or additional details. |  |
| 5.4.2 | Selection of methods | The laboratory shall use test and/or calibration methods, including methods for sampling, which meet the needs of the customer and which are appropriate for the tests and/or calibrations it undertakes. Methods published in international, regional or national standards (latest valid edition - unless it is not appropriate or possible, and supplemented with necessary additional details to ensure consistent use) shall preferably be used. Appropriate methods published by reputable technical organizations, or in relevant scientific texts or journals, or as specified by the manufacturer of the equipment or methods developed or adopted by the laboratory may also be used if they are validated. The laboratory shall confirm that it can properly operate standard methods before introducing the tests or calibrations. If the standard method changes, the confirmation shall be repeated. The laboratory shall inform the customer when the method proposed by the customer is considered to be inappropriate or out of date. The customer shall be informed as to the method chosen. |  |
| 5.4.3 | Laboratory-developed methods | The introduction of test and calibration methods developed by the laboratory for its own use shall be a planned activity and shall be assigned to qualified personnel equipped with adequate resources. |  |
| 5.4.4 | Non-standard methods | When it is necessary to use methods not covered by standard methods, these shall have been validated appropriately before use. NOTE: Description of procedures should contain at least the following information: appropriate identification; scope; description of the type of item to be tested or calibrated; parameters or quantities and ranges to be determined; apparatus and equipment, including technical performance requirements; reference standards and reference materials required; environmental conditions required and any stabilization period needed; affixing of identification marks, handling, transporting, storing and preparation of items; checks to be made before the work is started; checks that the equipment is working properly and, where required, calibration and adjustment of the equipment before each use; the method of recording the observations and results; any safety measures to be observed; criteria and/or requirements for approval/rejection; data to be recorded and method of data analysis and presentation; the uncertainty or the procedure for estimating uncertainty. |  |
| 5.4.5 | Validation of methods | Validation is the confirmation by examination and the provision of objective evidence that the particular requirements for a specific intended use are fulfilled. The laboratory shall validate non-standard methods, laboratory-designed/developed methods, standard methods used outside their intended scope, and amplifications and modifications of standard methods to confirm that the methods are fit for the intended use. The validation shall be as extensive as is necessary to meet the needs of the given application or field of application. The laboratory shall record the results obtained, the procedure used for the validation, and a statement as to whether the method is fit for the intended use. The range and accuracy of the values obtainable from validated methods (e.g. the uncertainty of the results, detection limit, selectivity of the method, linearity, limit of repeatability and/or reproducibility, robustness against external influences and/or cross-sensitivity against interference from the matrix of the sample/test object), as assessed for the intended use, shall be relevant to the customers' needs. NOTE: Validation is always a balance between costs, risks and technical possibilities. Validation includes specification of the requirements, determination of the characteristics of the methods, a check that the requirements can be fulfilled by using the method and a statement on the validity. The techniques used for the determination of the performance of a method should be one of, or a combination of, the following: calibration using reference standards or reference materials; comparison of results achieved with other methods; inter-laboratory comparisons; systematic assessment of the factors influencing the result; assessment of the uncertainty of the results based on scientific understanding of the theoretical principles of the method and practical experience. |  |
| No. | Title of the clause | Extracts from text of 17025:2005 and key words (arranged in blocks) | Comment |
| 5.4.6 | Estimation of uncertainty of measurement | A calibration laboratory, or a testing laboratory performing its own calibrations, shall have and shall apply a procedure to estimate the uncertainty of measurement for all calibrations and types of calibrations. All uncertainty components which are of importance in the given situation shall be taken into account using appropriate methods of analysis.If the nature of the test method precludes rigorous, metrologically and statistically valid, calculation of uncertainty of measurement the laboratory shall at least attempt to identify the components of uncertainty and make a reasonable estimation, and shall ensure that the form of reporting of the result does not give a wrong impression of the uncertainty. NOTE: In those cases where a well-recognized test method specifies limits to the values of the major sources of uncertainty of measurement and specifies the form of presentation of calculated results, the laboratory is considered to have satisfied this clause by following the test method and reporting instructions (see 5.10). |  |
| 5.4.7 | Control of data | Calculations and data transfers shall be subject to appropriate checks in a systematic manner. When computers or automated equipment are used for the acquisition, processing, recording, reporting, storage or retrieval of test or calibration data, the laboratory shall ensure that: computer software developed by the user is documented in sufficient detail and is suitably validated as being adequate for use; procedures are established and implemented for protecting the data; such procedures shall include, but not be limited to, integrity and confidentiality of data entry or collection, data storage, data transmission and data processing; computers and automated equipment are maintained to ensure proper functioning and are provided with the environmental and operating conditions necessary to maintain the integrity of test and calibration data. NOTE: Commercial off-the-shelf software (e.g. wordprocessing, database and statistical programmes) in general use within their designed application range may be considered to be sufficiently validated. However, laboratory software configuration/modifications should be validated. |  |
| 5.5 | Equipment | The laboratory shall be furnished with all items of sampling, measurement and test equipment required for the correct performance of the tests and/or calibrations and for achieving the accuracy required. Equipment shall comply with specifications relevant to the tests and/or calibrations concerned. Equipment shall be operated by authorized personnel having readily available up-to-date instructions on the use and maintenance of equipment (including any relevant manuals). The laboratory shall have procedures for safe handling, transport, storage, use and planned maintenance of measuring equipment to ensure proper functioning and in order to prevent contamination or deterioration. Equipment that has been subjected to overloading or mishandling, gives suspect results, or has been shown to be defective or outside specified limits, shall be taken out of service. It shall be isolated to prevent its use or clearly labelled or marked as being out of service until it has been repaired and shown by calibration or test to perform correctly. The laboratory shall examine the effect of the defect or departure from specified limits on previous tests and/or calibrations and shall institute the “Control of nonconforming work” procedure (see 4.9). Each item of equipment and its software used for testing and calibration and significant to the result shall, when practicable, be uniquely identified. Records shall be maintained of each item of equipment and its software significant to the tests and/or calibrations performed. The records shall include at least the following: the identity of the item of equipment and its software; the manufacturer's name, type identification, and serial number or other unique identification; checks that equipment complies with the specification (see 5.5.2); the current location, where appropriate; the manufacturer's instructions, if available, or reference to their location; dates, results and copies of reports and certificates of all calibrations, adjustments, acceptance criteria, and the due date of next calibration; the maintenance plan, where appropriate, and maintenance carried out to date; any damage, malfunction, modification or repair to the equipment. Calibration programmes shall be established for key quantities or values of the instruments. Before being placed into service or when intermediate checks are needed, equipment shall be calibrated or checked according to a defined procedure to establish that it meets the relevant specifications. Whenever practicable, all equipment requiring calibration shall be labelled, or otherwise identified to indicate the status of calibration, the date when last calibrated and the date or criteria when recalibration is due. Test and calibration equipment, including both hardware and software, shall be safeguarded from adjustments which would invalidate the test and/or calibration results. |  |
| No. | Title of the clause | Extracts from text of 17025:2005 and key words (arranged in blocks) | Comment |
| 5.6 | Measurement traceability |  |  |
| 5.6.1 | General | The laboratory shall have an established programme and procedure for the calibration of its equipment. |  |
| 5.6.2 | Specific requirements |  |  |
| 5.6.2.1 | Calibration | For calibration laboratories, the programme for calibration of equipment shall be designed and operated so as to ensure that calibrations and measurements made by the laboratory are traceable to the International System of Units (SI) (Système international d'unités) by means of an unbroken chain of calibrations or comparisons linking them to relevant primary standards of the SI units of measurement. The link to SI units may be achieved by reference to national measurement standards (primary or agreed representations of SI units or secondary which are standards calibrated by another national metrology institute). When using external calibration services, traceability of measurement shall be assured by the use of calibration services from laboratories that can demonstrate competence, measurement capability and traceability. The calibration certificates issued by these laboratories shall contain the measurement results, including the measurement uncertainty and/or a statement of compliance with an identified metrological specification (see also 5.10.4.2). There are certain calibrations that currently cannot be strictly made in SI units. In these cases calibration shall provide confidence in measurements by establishing traceability to appropriate measurement standards such as: the use of certified reference materials provided by a competent supplier to give a reliable physical or chemical characterization of a material; the use of specified methods and/or consensus standards that are clearly described and agreed by all parties concerned. Participation in a suitable programme of interlaboratory comparisons is required where possible. |  |
| 5.6.2.2 | Testing | For testing laboratories, the requirements given in 5.6.2.1 apply for measuring and test equipment with measuring functions used, unless it has been established that the associated contribution from the calibration contributes little to the total uncertainty of the test result. When this situation arises, the laboratory shall ensure that the equipment used can provide the uncertainty of measurement needed. Where traceability of measurements to SI units is not possible and/or not relevant, the same requirements for traceability to, for example, certified reference materials, agreed methods and/or consensus standards, are required as for calibration laboratories. |  |
| 5.6.3 | Reference standards and reference materials |  |  |
| 5.6.3.1 | Reference standards | The laboratory shall have a programme and procedure for the calibration of its reference standards. Reference standards shall be calibrated by a body that can provide traceability as described in 5.6.2.1. Such reference standards of measurement held by the laboratory shall be used for calibration only and for no other purpose, unless it can be shown that their performance as reference standards would not be invalidated. Reference standards shall be calibrated before and after any adjustment. |  |
| 5.6.3.2 | Reference materials | Reference materials shall, where possible, be traceable to SI units of measurement, or to certified reference materials. Internal reference materials shall be checked as far as is technically and economically practicable. |  |
| 5.6.3.3 | Intermediate checks | Checks needed to maintain confidence in the calibration status of reference, primary, transfer or working standards and reference materials shall be carried out according to defined procedures and schedules. |  |
| 5.6.3.4 | Transport and storage | The laboratory shall have procedures for safe handling, transport, storage and use of reference standards and reference materials in order to prevent contamination or deterioration and in order to protect their integrity. |  |
| 5.7 | Sampling | The laboratory shall have a sampling plan and procedures for sampling and for recording relevant data and operations relating to sampling when it carries out sampling of substances, materials or products for subsequent testing or calibration. The sampling plan as well as the sampling procedure shall be available at the location where sampling is undertaken. Sampling plans shall, whenever reasonable, be based on appropriate statistical methods. The records shall include the sampling procedure, the identification of the sampler, environmental conditions (if relevant) and diagrams or other equivalent means to identify the sampling location as necessary and, if appropriate, the statistics the sampling procedures are based upon. NOTE: Sampling is a defined procedure whereby a part of a substance, material or product is taken to provide for testing or calibration of a representative sample of the whole. |  |
| No. | Title of the clause | Extracts from text of 17025:2005 and key words (arranged in blocks) | Comment |
| 5.8 | Handling of test and calibration items | The laboratory shall have procedures for the transportation, receipt, handling, protection, storage, retention and/or disposal of test and/or calibration items, including all provisions necessary to protect the integrity of the test or calibration item, and to protect the interests of the laboratory and the customer A sampling procedure and information on storage and transport of samples, including information on sampling factors influencing the test or calibration result, should be provided to those responsible for taking and transporting the samples. It shall also have a system for identifying test and/or calibration items. The identification shall be retained throughout the life of the item in the laboratory. The system shall be designed and operated so as to ensure that items cannot be confused physically or when referred to in records or other documents. Upon receipt of the test or calibration item, abnormalities or departures from normal or specified conditions, as described in the test or calibration method, shall be recorded. When there is doubt as to the suitability of an item for test or calibration, or when an item does not conform to the description provided, or the test or calibration required is not specified in sufficient detail, the laboratory shall consult the customer for further instructions before proceeding and shall record the discussion. The laboratory shall have procedures and appropriate facilities for avoiding deterioration, loss or damage to the test or calibration item during storage, handling and preparation. When items have to be stored or conditioned under specified environmental conditions, or if a portion of an item is to be held secure (e.g., to enable complementary tests and/or calibrations to be performed later), these conditions shall be maintained, monitored and recorded. |  |
| 5.9 | Assuring the quality of test and calibration results | The laboratory shall have quality control procedures for monitoring the validity of tests and calibrations undertaken. The resulting data shall be recorded in such a way that trends are detectable and, where practicable, statistical techniques shall be applied to the reviewing of the results. Quality control data shall be analysed and, where they are found to be outside pre-defined criteria, planned action shall be taken to correct the problem and to prevent incorrect results from being reported. This monitoring shall be planned and reviewed and should be appropriate for the type and volume of the work undertaken. It may include, but not be limited to: regular use of certified reference materials and/or internal quality control using secondary reference materials; participation in interlaboratory comparison or proficiency-testing programmes; replicate tests or calibrations using the same or different methods; retesting or recalibration of retained items; correlation of results for different characteristics of an item. |  |
| 5.10 | Reporting the results | The results of each test, calibration, or series of tests or calibrations carried out by the laboratory shall be reported accurately, clearly, unambiguously and objectively, and in accordance with any specific instructions in the test or calibration methods. The results shall be reported, usually in a test report or a calibration certificate (issued as hard copy or by electronic data transfer provided that the requirements of ISO 17025:2005are met) and shall include all the information requested by the customer and necessary for the interpretation of the test or calibration results and all information required by the method used. This information is normally that required by 5.10.2, and 5.10.3 or 5.10.4. Any information listed in 5.10.2 to 5.10.4 which is not reported to the customer shall be readily available in the laboratory which carried out the tests and/or calibrations. |  |
| 5.10.1 | General |  |  |
| 5.10.2 | Test reports and calibration certificates | Each test report or calibration certificate shall include at least the following information: a title (e.g. “Test Report” or “Calibration Certificate”); the name and address of the laboratory, and the location where the tests and/or calibrations were carried out, if different from the address of the laboratory; unique identification of the test report or calibration certificate (such as the serial number), and on each page an identification, and a clear identification of the end; the name and address of the customer; reference to the sampling plan and procedures used by the laboratory or other bodies; identification of the method used; a description of, the condition of, and unambiguous identification of the item(s) tested or calibrated; the dates of receipt and of performance of the test or calibration of the test or calibration item(s); the test or calibration results with, where appropriate, the units of measurement; the name(s), function(s) and signature(s) or equivalent identification of person(s) authorizing the test report or calibration certificate; where relevant, a statement to the effect that the results relate only to the items tested or calibrated. |  |
| No. | Title of the clause | Extracts from text of 17025:2005 and key words (arranged in blocks) | Comment |
| 5.10.3 | Test reports | In addition to the requirements listed in 5.10.2, test reports shall, where necessary for the interpretation of the test results, include the following: deviations from, additions to, or exclusions from the test method, and information on specific test conditions, such as environmental conditions; where relevant, a statement of compliance/non-compliance with requirements and/or specifications; where applicable, a statement on the estimated uncertainty of measurement; where appropriate and needed, opinions and interpretations (see 5.10.5); additional information which may be required by specific methods, customers or groups of customers. Test reports containing the results of sampling shall include the following: the date of sampling; unambiguous identification of the substance, material or product sampled; the location of sampling, including any diagrams, sketches or photographs; a reference to the sampling plan and procedures used; details of any conditions during sampling that may affect the interpretation of the test results; any standard or other specification for the sampling method or procedure, and deviations, additions to or exclusions from the specification concerned. |  |
| 5.10.4 | Calibration certificates | In addition to the requirements listed in 5.10.2, calibration certificates shall, where necessary for the interpretation of calibration results, include the following: the conditions under which the calibrations were made that have an influence on the measurement results; the uncertainty of measurement and/or a statement of compliance with an identified metrological specification or clauses thereof; evidence that the measurements are traceable. The calibration certificate shall relate only to quantities and the results of functional tests. If a statement of compliance with a specification is made, this shall identify which clauses of the specification are met or not met taking into account the uncertainty of measurement. When a statement of compliance with a specification is made omitting the measurement results and associated uncertainties, the laboratory shall record those results and maintain them for possible future reference. When an instrument for calibration has been adjusted or repaired, the calibration results before and after adjustment or repair, if available, shall be reported. |  |
| 5.10.5 | Opinions and interpretations | When opinions and interpretations are included, the laboratory shall document the basis upon which the opinions and interpretations have been made. Opinions and interpretations shall be clearly marked as such in a test report. |  |
| 5.10.6 | Testing and calibration results obtained from subcontractors | When the test report contains results of tests performed by subcontractors, these results shall be clearly identified. The subcontractor shall report the results in writing or electronically. When a calibration has been subcontracted, the laboratory performing the work shall issue the calibration certificate to the contracting laboratory. |  |
| 5.10.7 | Electronic transmission of results | In the case of transmission of test or calibration results by telephone, telex, facsimile or other electronic or electromagnetic means, the requirements of ISO 17025:2005 shall be met. (see also 5.4.7). |  |
| 5.10.8 | Format of reports and certificates | The format shall be designed to accommodate each type of test or calibration carried out and to minimize the possibility of misunderstanding or misuse. |  |
| 5.10.9 | Amendments to test reports and calibration certificates | Material amendments to a test report or calibration certificate after issue shall be made only in the form of a further document, or data transfer, which includes the statement: “Supplement to Test Report [or Calibration Certificate], serial number... [or as otherwise identified]”, or an equivalent form of wording. Such amendments shall meet all the requirements of ISO 17025:2005. When it is necessary to issue a complete new test report or calibration certificate, this shall be uniquely identified and shall contain a reference to the original that it replaces. |  |
| Annex A | Nominal cross-references to ISO 9001:2000 |  |  |
| Annex B | Guidelines for establishing applications for specific fields |  |  |
|  | Bibliography |  |  |

# Footnotes

The structure of the first two columns follows the structure of the International Standard ISO/IEC 17025:2005 standard. However, text at lower structural levels is arranged into logical blocks in the third column and does not always reflect the structural order of the International Standard. For the full wording of the Norm it is strongly recommended to use this table together with the original document.

The term 'management system' in ISO 17025:2005 means the quality, administrative and technical systems that govern the operations of a laboratory.